## Instructions
Instructions
You have 3 sheets in this file: 1) The instructions, 2) Enter data here, 3) EXAMPLE
!!! Only fill the information into the grey cells !!!
1. Read carefully the User Manual of the iCheck device you will be operating.
2. Turn on your iCheck and get acquainted with the menu structure and functions.
3. Prepare one sample of your vitamin A fortified target food (i.e. sugar, wheat flour) by diluting it with water so that expected concentration of diluted sample was within iCheck Fluoro measurement range of 50 to 3000 µg RE/L [167 - 10,000 IU/kg].
4. Inject and measure the diluted sample minimum 5 times. Make sure that the extraction time for each injection is at least 5 minutes. Make sure that you SHAKE vigorously the reagent vial right after sample injection and after 5 minutes. Make sure that there is a clear phase separation, and the upper phase is clear.
5. Record the results in this training sheet. Your average result, standard deviation and coefficient of variation (CV) are automatically calculated.
6. Coefficient of variation of your results for must be below 9% for vitamin premix and sauces, below 15% for flour, below 3% for milk, and below 5% for sugar.
7. If your CV is higher than 9% for vitamin premix and sauces, higher than 15% for flour, higher than 3% for milk, and higher than 5% for sugar, carefully read the section Instructions in the User Manual, mix well your diluted sample and repeat the measurement another 5 times.
8. Flour and other food types (KNOWN exceptions: milk, breast milk, vitamin premix, standard) interfer with iCheck Fluoro measurement due to matrix effect. It is recommended to measure following identical sample preparation protocol the unfortified food sample to estimate the level of the interference and correct the result accordingly. Example, average matrix effect of wheat flour is 0.65 mg/kg. This means that unfortified flour will give false positive result of on average 0.65 mg/kg. Deduct 0.65 mg/kg from your result with fortified sample.

## Enter data here_print
Practical Training Protocol
Vitamin A Analysis in Food
Name of the operator
Date
Sample description
Expected vitamin A concentration in the sample [mg/kg = ppm]
Water [g] = [mL]
Sample [g]
Total Diluted Sample Volume [mL]
Dilution Factor (Total Diluted Sample Volume [mL] /Sample [g] )
iCheck FLUORO reagent vial LOT
iCheck FLUORO DEVICE ID No. (on the back of the device)
FLUORO Standard LOT No.
FLUORO Standard reading on your iCheck [mg/L] (use DEVICE CONTROL mode)
|  | Injection time [hh:mm] | Measurement time [hh:mm] | iCheck Result [µg RE/L]
iCheck™ Flouro measurement result | Vial 1
Vial 2
Vial 3
Vial 4
Vial 5
Vial 6
Vial 7
Vial 8
Vial 9
Vial 10
Average iCheck Result [µg RE/L]
Standard Deviation [µg RE/L]
Coefficient of Variation (CV) %
Vitamin A in the sample before dilution (Average Result x Dilution Factor / 1000) [ppm=mg/kg]
Estimated Matrix effect of your sample in ppm (mg/kg)
Vitamin A in sugar, [ppm=mg/kg]

## Enter data here
Practical Training Protocol
Vitamin A Analysis in Food
Name of the operator
Date
Sample description
Expected vitamin A concentration in the sample [mg/kg = ppm]
Water [g] = [mL]
Sample [g]
Total Diluted Sample Volume [mL]
Dilution Factor (Total Diluted Sample Volume [mL] /Sample [g] ) |  | #DIV/0!
Expected vitamin A concentration in the DILUTED sample [µg RE/L] |  | #DIV/0!
iCheck Fluoro reagent vial LOT (batch)
iCheck Fluoro device ID No. (FL xxx)
Fluoro Standard LOT (old Calibration Control)
Fluoro Standard reading, µg RE/L
|  | Injection time [hh:mm] | Measurement time [hh:mm] | iCheck Result [µg RE/L]
iCheck™ Flouro measurement result | Vial 1
Vial 2
Vial 3
Vial 4
Vial 5
Vial 6
Vial 7
Vial 8
Vial 9
Vial 10
Average iCheck Result [µg RE/L] |  |  |  | #DIV/0!
Standard Deviation [µg RE/L] |  |  |  | #DIV/0!
Coefficient of Variation (CV) % |  |  |  | #DIV/0!
Vitamin A in the sample before dilution (Average Result x Dilution Factor x 1000) [ppm=mg/kg] |  |  |  | #DIV/0!
Estimated Matrix effect of your sample in ppm (mg/kg)
Vitamin A in flour sample after correction for the matrix effect, [ppm=mg/kg] |  |  |  | #DIV/0!

## EXAMPLE
Practical Training Protocol
Vitamin A Analysis in Food
Name of the operator |  | John Smith
Date |  | 2014-10-10
Sample description |  | Flour collected from the open market on the Green Street in Cairo on 09.20.2014. Brand "Wheat", Producer "XXX"
Expected vitamin A concentration in the sample [mg/kg = ppm] |  | 3
Water [g] = [mL] |  | 100
Sample [g] |  | 20
Total Diluted Sample Volume [mL] |  | 114
Dilution Factor (Total Diluted Sample Volume [mL] /Sample [g] ) |  | 5.7
Expected vitamin A concentration in the DILUTED sample [µg RE/L] |  | 526.3158
iCheck Fluoro reagent vial LOT (batch) |  | LOT 2 - 05.2014
iCheck Fluoro device ID No. (FL xxx) |  | FL 22233344
Fluoro Standard LOT (old Calibration Control) |  | LOT 2 - 05.2014
Fluoro Standard reading, µg RE/L |  | 1800
|  | Injection time [hh:mm] | Measurement time [hh:mm] | iCheck Result [µg RE/L]
iCheck™ Flouro measurement result | Vial 1 | 10:10 | 10:15 | 530
| Vial 2 | 10:12 | 10:17 | 590
| Vial 3 | 10:14 | 10:19 | 620
| Vial 4 | 10:16 | 10:21 | 530
| Vial 5 | 10:18 | 10:23 | 690
| Vial 6 | 10:20 | 10:25 | 660
| Vial 7 | 10:22 | 10:27 | 587
| Vial 8 | 10:24 | 10:29 | 688
| Vial 9 | 10:26 | 10:31 | 700
| Vial 10 | 10:28 | 10:33 | 613
Average iCheck Result [µg RE/L] |  |  |  | 620.8
Standard Deviation [µg RE/L] |  |  |  | 63.0305
Coefficient of Variation (CV) % |  |  |  | 10.15%
Vitamin A in the sample before dilution (Average Result x Dilution Factor x 1000) [ppm=mg/kg] |  |  |  | 3.5386
Estimated Matrix effect of your sample in ppm (mg/kg) |  |  |  | 0.65
Vitamin A in flour sample after correction for the matrix effect, [ppm=mg/kg] |  |  |  | 2.8886